FACSort /calibur Sorter flush

- Turn off the sorter.

- Use 60-mL syringe full of DI water and bleed out any air. Remove sheath filter and connect adaptor from syringe to top sheath filter port.

- Place tube of DI water on sip, set fluid control to RUN.

- Install 60-mL tube in first collection port of sorter and press the sort-line purge button.

- Depress the syringe plunger until approximately 10-mL of fluid has collected into the first port. Repeat process for 2nd  and  3rd collection port.

- (ONLY if option is present) Turn on the cell concentrator module by pressing the power button, press the sort-line purge button and flush 20-mL of DI water as line to concentrator is twice as long as the first two. Turn off cell concentrator module.

- Replace sheath filter.  Drain and fill the flow cell to remove any air bubbles.