## Slide 1: Flow Cytometry and Sorting Part 1
Lecture Notes for “Fluorescence Spectroscopy in Biological Research”
Robert F. Murphy, October 1996

## Slide 2: Sources
Flow Cytometry and Sorting, 2nd ed. (M.R. Melamed, T. Lindmo, M.L. Mendelsohn, eds.), Wiley-Liss, New York, 1990 - referred to here as MLM
Flow Cytometry: Instrumentation and Data Analysis (M.A. Van Dilla, P.N. Dean, O.D. Laerum, M.R. Melamed, eds.), Academic Press, London, 1985 - VDLM

## Slide 3: Sources (continued)
The Purdue
Cytometry
CD ROM
Volume 1 - 1996
Purdue University Cytometry Laboratories
Home Page | Table of Contents | Sponsors | Sample WEB Pages

## Slide 4: Definitions
Flow Cytometry
  Measuring properties of cells in flow
Flow Sorting
  Sorting (separating) cells based on properties measured in flow
  Also called Fluorescence-Activated Cell Sorting (FACS)

## Slide 5: Basics of Flow Cytometry
Cells in suspension
flow in single-file through
an illuminated volume where they
scatter light and emit fluorescence
that is collected, filtered and
converted to digital values
that are stored on a computer
Fluidics
Optics
Electronics

## Slide 6: Fluidics
Need to have cells in suspension flow in single file through an illuminated volume
In most instruments, accomplished by injecting sample into a sheath fluid as it passes through a small (50-300 µm) orifice

## Slide 7
Flow Cell
Injector
Tip
Fluorescence
signals
Focused laser
beam
Sheath
fluid
Purdue University Cytometry Laboratories

## Slide 8: Fluidics
When conditions are right, sample fluid flows in a central core that does not mix with the sheath fluid
This is termed Laminar flow

## Slide 9: Fluidics - Laminar Flow
Whether flow will be laminar can be determined from the Reynolds number
When Re < 2300, flow is always laminar
When Re > 2300, flow can be turbulent

## Slide 10: Fluidics
The introduction of a large volume into a small volume in such a way that it becomes “focused” along an axis is called Hydrodynamic Focusing

## Slide 11: Fluidics
The figure shows the mapping between the flow lines outside and inside of a narrow tube as fluid undergoes laminar flow (from left to right). The fluid passing through cross section A outside the tube is focused to cross section a inside.
V. Kachel, H. Fellner-Feldegg & E. Menke - MLM Chapt. 3

## Slide 12: Fluidics
V. Kachel, H. Fellner-Feldegg & E. Menke - MLM Chapt. 3
Notice how the ink is focused into a tight stream as it is drawn into the tube under laminar flow conditions.
Notice also how the position of the inner ink stream is influenced by the position of the ink source.

## Slide 13: Fluidics
V. Kachel, H. Fellner-Feldegg & E. Menke - MLM Chapt. 3
Notice how the ink is focused into a tight stream as it is drawn into the tube under laminar flow conditions.
Notice also how the position of the inner ink stream is influenced by the position of the ink source.

## Slide 14: Fluidics
How do we accomplish sample injection and regulate sample flow rate?
  Differential pressure
  Volumetric injection

## Slide 15: Fluidics - Differential Pressure System
Use air (or other gas) to pressurize sample and sheath containers
Use pressure regulators to control pressure on each container separately

## Slide 16: Fluidics - Differential Pressure System
Sheath pressure will set the sheath volume flow rate (assuming sample flow is negligible)
Difference in pressure between sample and sheath will control sample volume flow rate
Control is not absolute - changes in friction cause changes in sample volume flow rate

## Slide 17: Fluidics - Differential Pressure System
C. Göttlinger, B. Mechtold, and A. Radbruch

## Slide 18: Fluidics - Volumetric Injection System
Use air (or other gas) pressure to set sheath volume flow rate
Use syringe pump (motor connected to piston of syringe) to inject sample
Sample volume flow rate can be changed by changing speed of motor
Control is absolute (under normal conditions)

## Slide 19: Fluidics - Volumetric Injection System
H.B. Steen - MLM Chapt. 2

## Slide 20: Fluidics - Particle Orientation and Deformation
As cells (or other particles) are hydrodynamically focused, they experience different shear stresses on different points on their surfaces (an in different locations in the stream)
These cause cells to orient with their long axis (if any) along the axis of flow

## Slide 21: Fluidics - Particle Orientation and Deformation
The shear stresses can also cause cells to deform (e.g., become more cigar-shaped)

## Slide 22: Fluidics - Particle Orientation and Deformation
“a: Native human erythrocytes near the margin of the core stream of a short tube (orifice). The cells are uniformly oriented and elongated by the hydrodynamic forces of the inlet flow.
b: In the turbulent flow near the tube wall, the cells are deformed and disoriented in a very individual way. v>3 m/s.”
V. Kachel, et al. - MLM Chapt. 3

## Slide 23: Fluidics - Flow Chambers
The flow chamber
  defines the axis and dimensions of sheath and sample flow
  defines the point of optimal hydrodynamic focusing
  can also serve as the interrogation point (the illumination volume)

## Slide 24: Fluidics - Flow Chambers
Four basic flow chamber types
  Jet-in-air
    best for sorting, inferior optical properties
  Flow-through cuvette
    excellent optical properties, can be used for sorting
  Closed cross flow
    best optical properties, can’t sort
  Open flow across surface
    best optical properties, can’t sort

## Slide 25: Fluidics - Flow Chambers
H.B. Steen - MLM Chapt. 2
Jet-in-air nozzle (sense in air)

## Slide 26: Fluidics - Flow Chambers
H.B. Steen - MLM Chapt. 2
Flow through cuvette (sense in quartz)

## Slide 27: Fluidics - Flow Chambers
H.B. Steen - MLM Chapt. 2
Closed cross flow chamber

## Slide 28: Optics
Need to have a light source focused on the same point where cells have been focused (the illumination volume)
Two types of light sources
  Lasers
  Arc-lamps

## Slide 29: Optics - Light Sources
Lasers
  can provide a single wavelength of light (a laser line) or (more rarely) a mixture of wavelengths
  can provide from milliwatts to watts of light
  can be inexpensive, air-cooled units or expensive, water-cooled units
  provide coherent light

## Slide 30: Optics - Light Sources
Arc-lamps
  provide mixture of wavelengths that must be filtered to select desired wavelengths
  provide milliwatts of light
  inexpensive, air-cooled units
  provide incoherent light

## Slide 31: Optics - Optical Channels
An optical channel is a path that light can follow from the illuminated volume to a detector
Optical elements provide separation of channels and wavelength selection

## Slide 32: Optics - Forward Scatter Channel
When a laser light source is used, the amount of light scattered in the forward direction (along the same axis that the laser light is traveling) is detected in the forward scatter channel
The intensity of forward scatter is proportional to the size, shape and optical homogeneity of cells (or other particles)

## Slide 33
Forward Angle Light Scatter
FALS Sensor
Laser
Purdue University Cytometry Laboratories

## Slide 34: Optics - Side Scatter Channel
When a laser light source is used, the amount of light scattered to the side (perpendicular to the axis that the laser light is traveling) is detected in the side or 90o scatter channel
The intensity of side scatter is proportional to the size, shape and optical homogeneity of cells (or other particles)

## Slide 35
90 Degree Light Scatter
FALS Sensor
90LS Sensor
Laser
Purdue University Cytometry Laboratories

## Slide 36: Optics - Light Scatter
Forward scatter tends to be more sensitive to surface properties of particles (e.g., cell ruffling) than side scatter
  can be used to distinguish live from dead cells
Side scatter tends to be more sensitive to inclusions within cells than forward scatter
  can be used to distinguish granulated cells from non-granulated cells

## Slide 37: Optics - Fluorescence Channels
The fluorescence emitted by each fluorochrome is usually detected in a unique fluorescence channel
The specificity of detection is controlled by the wavelength selectivity of optical filters and mirrors

## Slide 38
Laser
Fluorescence Detectors
Fluorescence
FALS Sensor
Fluorescence detector
(PMT3, PMT4 etc.)
Purdue University Cytometry Laboratories

## Slide 39: Optics - Filter Properties
Optical filters are constructed from materials that absorb certain wavelengths (while transmitting others)
Transitions between absorbance and transmission are not perfect; the sharpness can be specified during filter design

## Slide 40: Optics - Filter Properties
When using laser light sources, filters must have very sharp cutons and cutoffs since there will be many orders of magnitude more scattered laser light than fluorescence
Can specify wavelengths that filter must reject to certain tolerance (e.g., reject 488 nm light at 10-6 level: only 0.0001% of incident light at 488 nm gets through)

## Slide 41: Optics - Filter Properties
Long pass filters transmit wavelengths above a cut-on wavelength
Short pass filters transmit wavelengths below a cut-off wavelength
Band pass filters transmit wavelengths in a narrow range around a specified wavelength
  Band width can be specified

## Slide 42: Standard Long Pass Filters
Transmitted Light
Light Source
520 nm Long Pass Filter
>520 nm Light
Transmitted Light
Light Source
575 nm Short Pass Filter
<575 nm Light
Standard Short Pass Filters
Purdue University Cytometry Laboratories

## Slide 43
Standard Band Pass Filters
Transmitted Light
White Light Source
630 nm BandPass Filter
620 -640 nm Light
Purdue University Cytometry Laboratories

## Slide 44: Optics - Filter Properties
When a filter is placed at a 45o angle to a light source, light which would have been transmitted by that filter is still transmitted but light that would have been blocked is reflected (at a 90o angle)
Used this way, a filter is called a dichroic filter or dichroic mirror

## Slide 45: Dichroic Filter/Mirror
Filter placed at 45o
Reflected light
Transmitted Light
Light Source
original from Purdue University Cytometry Laboratories; modified by R.F. Murphy

## Slide 46: Optics - Filter Layout
To simultaneously measure more than one scatter or fluorescence from each cell, we typically use multiple channels (multiple detectors)
Design of multiple channel layout must consider
  spectral properties of fluorochromes being used
  proper order of filters and mirrors

## Slide 47
Ethidium
PE
cis-Parinaric acid
Texas Red
PE-TR Conj.
PI
FITC
600 nm
300 nm
500 nm
700 nm
400 nm
457
350
514
610
632
488
Common
Laser
Lines
Purdue University Cytometry Laboratories

## Slide 48
PMT
PMT
PMT
PMT
Dichroic
Filters
Bandpass
Filters
Example Channel Layout for Laser-based Flow Cytometry
Laser
1
2
3
4
Flow cell
original from Purdue University Cytometry Laboratories; modified by R.F. Murphy

## Slide 49: Example Channel Layout for Arc Lamp-based Flow Cytometry
(Overhead 10)
H.B. Steen - MLM Chapt. 2

## Slide 50: Optics - Detectors
Two common detector types
  Photodiode
    used for strong signals when saturation is a potential problem (e.g., forward scatter detector)
  Photomultiplier tube (PMT)
    more sensitive than photodiode but can be destroyed by exposure to too much light

## Slide 51: Optics - Wavelength Dependence of Photomultipliers
H.B. Steen - MLM Chapt. 2
We should consider the properties of PMTs when designing an optical layout; knowledge of PMT types on a particular instrument allows optimum use of available fluorescence channels

## Slide 52: Summary of Part 1
Cells in suspension
flow in single-file through
an illuminated volume where they
scatter light and emit fluorescence
that is collected, filtered and
converted to digital values
that are stored on a computer
Fluidics
Optics
Electronics